MINISTRY OF HEALTH & FAMILY WELFARE

Government of India

Through

United Nations Office for Project Services, 11 Golf Links, New Delhi-110003

PROCUREMENT OF BLOOD MOBILES

IFB No.  UNOPS India/NACO/BM/21/ 2008

Amendment –II, dt. August 25, 2008

Following modification to the Bid Document is made through this Amendment for the procurement of Blood Mobiles, with reference to above IFB.

| (A) IFB & Bid data Sheet |  |  |
| --- | --- | --- |
| Wherever appearing in the bid document following modification is made |  |  |
| 1. | Last Date For Sale Of Bidding Document : 1200 Hrs On 02 -09-2008 Last Date And Time For Receipt Of Bid: 1400 hrs on 02-09-2008 Time And Date Of Opening Of Bids: 1415 hrs on 02-09-2008 | Last Date For Sale Of Bidding Document : 1200 Hrs On 10 -09-2008 Last Date And Time For Receipt Of Bid: 1400 hrs on 10-09-2008 Time And Date Of Opening Of Bids: 1415 hrs on 10-09-2008 |
| ITB 20.1 of BDS |  |  |
| The bid validity period shall be up to January 02, 2009 |  | The bid validity period shall be up to January 10, 2009. |

| 2. Technical Specifications |  |
| --- | --- |
| For | Read |
| Point No 2. |  |
| The engine should be commensurate with the requirements that should run both on diesel as well as on CNG and air-conditioning on Generator. | The engine should be commensurate with the requirements that should run on diesel and air-conditioning on Generator. |
| Point No. 6 |  |
| Dual electricity systems with facility for change over and cable for external power source | Dual electricity systems with facility for change over and cable for external power source having length of minimum of 50 meters. |
| Point No. 22 |  |
| Facility for GOS installation | Facility for GPS installation |
| Internal Designs :- Flooring |  |
| The entire vehicle should follow bio-safety norms. | The entire vehicle should follow bio-safety norms and flooring should be acid and alkali resistant and impervious to water. |
| Part 2, Section VI. Schedule of Requirements- Delivery Schedule |  |
| Within 90 days, from the date of signing of the Contract for Group A, B & C Bidders | Within 120 days from the date of signing of the contract for Group A, B & C Bidders. (60 days for developing the prototype and getting the approval and another 60 days for completing the delivery.) |
| Additional Requirements, Technical Specifications A. Donor Couch - Nos 04 1. Comfortable chair type with soft padding for cushioning and rexin cover. 2. Seat, back rest and leg rest size designed for donor comfort. Seat height approximately 58 – 60 cm. 3. Adjustable arm rest for donor’s comfort and phlebotomist friendly. 4. Easily tilted to head low position, electrically operated. 5. To be operational on 220 to 240 V at 50 Hz. B. Blood Transportation Boxes – Nos 04 1. Fixed two temperatures setting adjustable to +4°C and +22°C. 2. Capacity to accommodate minimum 48 blood bags of 450 ml. 3. External dimension of each transportation box should be within 1000 -1500mm (L)X 800-1000 mm (W)X 600- 700 mm (H) 4. Chest / Vertical Cabinet with built in handles and hinged door. 5. Body material should be Polyethylene (corrosion free) 6. Smooth castors with stabilizers for flexibility and movement. 7. Ambient working temperature -2°C to +43°C. 8. Main power switch. 9. Microprocessor based temperature controller with integrated audio visual temperature alarm function on digital monitoring display. 10. Separate selector switch for + 4°C and +22°C temperature settings. 11. Nominal DC power 12 / 24 volt and Nominal AC power 100 – 240 volt, 50/60Hz. 12. Automatic AC / DC power supply selection. 13. Automatic Cooling / heating operation. 14. Air cooled hermetically sealed DC compressor. 15. Vehicle fixation kit. 16. Controlled fan cooling system for constant air/temperature distribution. 17. For storage of Blood bags, minimum two (2) wire basket should be provided in each box. 18. Power alarm audio and visual by LED. Audio alarm can be switched off. Temperature audio and visual high and low alarm by LED. - Minimum Polyurethane insulation CFC free for the cabinet and the door. C. Di-Electric Tube Sealer - Nos 02 1. Should be a heavy duty tube-sealer. 2. Should be for bench-top use. 3. The sealing time should not be more than 2 seconds. 4. Sealing triggering should be automatic. 5. Should also have extended portable hand unit. 6. Should have indication lamps. 7. No warm-up time should be required. 8. Should ensure easy separation of tube segments after the sealing. 9. Should be simple to handle. 10. To be operational on 220 to 240 V at 50 Hz. |  |

D.     Blood Mixer and Collector- Nos 04

| 1.Power Supply | 230 V + 10 ; 50Hz. |
| --- | --- |
| 2. Volume setting: | Pre-selection of volume to be collected. Tarring of bag volume before collection. Tarring range: 0 – 600 g. Automatic storage and recall of set volume. Measure volume with best accuracy <1%. |
| 3. Indications & alarms: | (i) LED indication on commencement of collection. (ii) LED indication and audible alarm at the end of collection. (iii) Iindication of time taken for collection. (iv) Indication of blood flow with audio alarm when blood flow is higher or lower than desired. (v) Continuous display of collected volume, flow and time during collection. |
| 4. Automatic clamping at termination of preset volume collection. |  |
| 5. Automatic release of bag when lifted. |  |
| 6. Continuous agitation of blood bags during collection: 12 – 16 rpm. |  |
| 7. Easy provision to change preset volume. |  |
| 8. Should be suitable for all types of bags. |  |

All other terms and conditions of the bid document remain unaltered.